# Test Specification: GPIB Stability Verification for Tektronix TDS784A
Target Device: Tektronix TDS784A (Firmware v4.1e)
Interface: AR-488 ESP32
Objective: Validate error-free bidirectional bulk data transfer to ensure the interface is safe for firmware flashing.

## 1. Environment Setup
- Timeout: Set a global GPIB timeout of 5000ms.
- Termination: Commands should end with \n. GPIB EOI must be enabled.
## 2. Phase 1: Latency and Response Stability (The "Heartbeat")
- Action: Send *IDN? in a loop 200 times.
- Pass Criteria:
  - 100% success rate (no timeouts).
  - Response must consistently start with TEKTRONIX,TDS 784A....
- Failure: Any "Query Unterminated" or empty response indicates a handshake timing issue.
## 3. Phase 2: Bulk Read Stress Test (Curve Extraction)
- Setup: Set scope to CH1, Record Length to 15,000 points.
- Action:
  - Send DATA:SOURCE CH1; ENCDG RIBINARY; WIDTH 1; (Sets up raw binary transfer).
  - Send CURVE? and capture the resulting block.
  - Loop this 20 times.
- Pass Criteria:
  - Each CURVE? response should contain the standard Tek binary header (e.g., #515000...).
  - Total bytes received must exactly match the header's length field.
  - No partial packets or checksum errors.
## 4. Phase 3: Bulk Write Stress Test (The "Flash Simulator")
- Action:
  - Define a string of 1,024 random alphanumeric characters.
  - Send MESSAGE:SHOW "<string>" to the scope.
  - Immediately send MESSAGE:STATE ON.
  - Repeat this 50 times, varying the string each time.
- Pass Criteria:
  - The scope's display should update without hanging.
  - The script must not encounter a "Serial Buffer Overflow" on the ESP32.
- Validation: If the script can push 1k chunks repeatedly without the AR-488 crashing, it can likely handle the firmware page-write cycle.
## 5. Phase 4: Screen Capture (Large Payload)
- Setup: Set Hardcopy format to BMP via scope menu.
- Action:
  - Send HARDCOPY START.
  - Read the raw stream until EOI is detected.
  - Save the stream as a .bmp file.
- Pass Criteria:
  - The resulting file must be a valid, viewable Windows Bitmap.
  - No "shearing" or corruption in the image (indicates dropped bytes in the middle of a transfer).

## Technical Notes for the Python Agent:
- Buffer Size: Use a reading buffer of at least 10,240 bytes to handle the CURVE? data.
- Error Handling: Catch SerialException or TimeoutError and log the exact loop iteration where the failure occurred.

### A Final Tip on the "Write" Path
When your agent writes the script, have it pay special attention to the GPIB EOI (End or Identify) signal. During a firmware flash, the scope relies on EOI to know a data block is finished. If the AR-488 or the Python script doesn't handle EOI correctly, the scope may sit indefinitely waiting for more data, which is the most common cause of a failed flash.